Supplement 1 – thoron decay correction

For air sampling through tubes from an exhalation container, 220Rn concentration at the instrument’s inlet, C1 (Bq m-3) can be expressed as (Durridge Company Inc 2023):

|  | C1= C0e-λtn*V1/q | (S.1) |
| --- | --- | --- |

where C0 (Bq m-3) is the concentration at sampling point (exhalation container), λtn is the decay constant for 220Rn (λtn = (12.42 ± 0.07) * 10-3 per second = (0.7452 ± 0.0042) per minute (Laraweb 2024)), V1 (L) is the tube volume and q (L/min) is the flow rate.

A tube with length 1.75 m and diameter of 3.5 mm has a volume of 0.017 L, and the ratio C0/C1 is,

|  | C0C1= e0.745*0.017/0,75=1.017 | (S.2) |
| --- | --- | --- |

For EQF we use recommended airflow of 0.75 L/min, but with calibration for 0.4 L/min, we correct for the equilibrium 220Rn concentration in the measurement chamber, as air fills the measurement chamber faster than the instrument was calibrated for. Thus, we solve the balance equation for 220Rn at equilibrium in the measurement chamber. Change of 220Rn particles N over time in the chamber of volume V2 (m3) is:

|  | dNdt= C1q-C2q-λN2=0 | (S.3) |
| --- | --- | --- |

where C2 (Bq m-3) is the equilibrium 220Rn concentration inside the measurement chamber at its outlet valve.  The last term is the decay of 220Rn in the chamber, where N2 = C2 * V2. Solving this for C2 gives us:

|  | C2= C11+λtn V2q | (S.4) |
| --- | --- | --- |

We use this to calculate the concentration in the chamber, C2, from C1,cal and qcal, and correct for the airflow rate, achieving a corrected 220Rn concentration at inlet valve C1,meas using qmeas = 0.75 L/min. According to the EQF 3220 manual, the internal instrument air loop volume is 250 mL. We assume this mainly comprise the chamber volume and set V2 = 0.25 L. The correction factor for using different flow speed is:

|  | C1,measC1,cal= 1+λtn V2qmeas1+λtnV2qcal =0.852 | (S.5) |
| --- | --- | --- |

Multiplying the two correction factors gives the final correction factor for 220Rn results using the external tube and 0.75 L/min flow rate, which equals 86.7 % of the recorded value.

Supplement 2 - Measurement Time

To ensure a low detection level, the reference measurement period for 220Rn was 30 minutes the first field day and 15 minutes for the rest. Generally, for exhalation series measurements, sampling time was reduced to attain sufficient measurements during  the linear phase. Measurement periods of 5, 2 and 1 minutes were tested to find the optimal counting statistics. For some 5-minute intervals, the linear phase had only two points and it was unclear whether both were actually part of the linear phase or the last was deviation so that exhalation rate was underestimated. Because of the high exhalation rate of 220Rn at the Mining Hill, even the 1-minute interval gave sufficient counting statistics (e.g Figure 1a main text). For 220Rn, all six methods for different accumulation phases to calculate exhalation rate can at this Fen locality be completed within an hour and offers a method to assess diurnal variation (but this does not hold for 222Rn). The fastest method is the linear phase methods, and this can give more rapid estimates of the diurnal variation. However, for methods including the threshold phase, 220Rn concentration fluctuations during threshold in some measurement series occurs and can be due to container leakage, wind effects, and variable solar heating of the exhalation container.

Supplement 3 – DEDUCTIONS FOR THE SIX METHODS

Method 1 – the old way, simulating sealed-cans

We use datapoints from the whole accumulation process to calculate the exhalation rate, Φ (Bq m-2 s-1), using the following equation.

|  | Φ= E∙V∙λS∙Teff | (S.6) |
| --- | --- | --- |

Where E (Bq m-3 s) is the 220Rn exposure, Teff=T+ 1λe-λT-1 (s) and T (s) is the duration of the measurement.

The uncertainty of the exhalation rate for method 1 (sealed can) is:

|  | uΦΦ=uEE2+uVV2+uλλ2+uSS2+uTeffTeff2 | (S.7) |
| --- | --- | --- |

different phases of accumulation

Change over time in activity concentration of 220Rn or 222Rn, C, in the exhalation container due to accumulation, can be described from the equation below. The first term is exhalation from the surface into the container. Due to surface irregularities, effective surface area is: S = 0.257 ± 0.013 m2. The second term is decay in the container where λi (s-1) is the is the decay constant of the radon isotope under investigation. The last two terms are leakage and back diffusion, where λL (s-1) is the time constant of leakage and λB (s-1) is the time constant of back diffusion.

|  | dCdt= Φ∙SV-λi∙C-λL∙C- λB∙C | (S.8) |
| --- | --- | --- |

Solving the equation shows the activity concentration of 220Rn can be described from the change in activity concentration over time (below), as known from ISO 11665-7:

|  | Ct= Φ∙SV∙λ1-e-λt | (S.9) |
| --- | --- | --- |

Where λ=λi+λL+λB.

thus follows Method 3-5 – Linear methods

Assuming the initial linear slope of the curve is independent of back diffusion, negligible leakage, and constant exhalation rate, the linear increase of activity concentration during accumulation is (simplified):

|  | Ct= Φ∙SV∙t | (S.13) |
| --- | --- | --- |

The exhalation rate can then be expressed as:

|  | Φ=A ∙VS | (S.14) |
| --- | --- | --- |

where A (Bq m-3 s-1) is the slope of the activity concentration as a function of time, which can be calculated using linear regression. The uncertainty of the exhalation rate calculated for the initial linear phase is:

|  | uΦΦ=uAA2+uVV2+uSS2 | (S.15) |
| --- | --- | --- |

Method 2 – threshold phase

Generally, when threshold is reached, the Rn concentration (Bq m-3) will be:

|  | Clim=Ct→∞= Φ∙SV∙λ | (S.10) |
| --- | --- | --- |

Hence, we can calculate the exhalation from the formula:

|  | Φ= Clim∙V∙λS | (S.11) |
| --- | --- | --- |

Uncertainty of the exhalation rate for this method is:

|  | uΦΦ=uClimClim2+uVV2+uλλ2+uSS2 | (S.12) |
| --- | --- | --- |

Method 6 – exponential fitting – using all phases of accumulation

The linear increase and the threshold phases are both included, using non-linear least squares exponential fitting (R Core Team 2016), and the model of Seo et al (2018).

|  | Ct= B1-e-kt | (S.16) |
| --- | --- | --- |

In addition, we include the reference measurement. Where B= Φ∙SV∙λ (Bq m-3) and k=λ=λi+λB+λL (s-1). Parameters B and k are used to estimate exhalation rate.

|  | Φ= B∙k∙VS | (S.17) |
| --- | --- | --- |

The uncertainty of the exhalation rate is:

|  | uΦΦ=uBB2+ukk2+uVV2+uSS2 | (S.18) |
| --- | --- | --- |

Supplement 3.6 – 222Rn grab sampling

Mixed atmospheres of 222Rn and 220Rn involve bias on measurements from 220Rn progeny affecting regions of interest (ROI) in alpha spectrometry where 222Rn progeny are counted (Turtiainen et al. 2022, Sarad AN-004). Since progeny of 220Rn are short-lived,  a grabbing procedure allowing its decay before measuring 222Rn, would give much lower detection limit and less uncertainty in measurements.

The half-life of 220Rn daughter 216Po is 0.148 s and secular equilibrium is instantaneous. Some alpha particles from 216Po lose energy when hitting detector walls and may be counted in the lower alpha energy region of interest (ROI) for 222Rn progeny 218Po. In Fen, where the 232Th series dominates, 220Rn thus poses a real problem for 222Rn measurements. We solved this by grab sampling into a closed tube system and letting 220Rn and 216Po decay before measuring 222Rn. First, we run air through the tubes for 5 minutes to fill it with the air to be sampled. To grab, we make a closed tube-loop between the instrument in-valve and the out-valve and wait for decay within the loop. Enough 220Rn needs to decay so that the tailing from 216Po does not affect the signal in the ROI for 218Po. In the redrock part of Fen complex, 220Rn has levels as high as 300 000 Bq/m3 in the exhalation container and therefore each waiting period after the grab-sample must last for at least 15 minutes (more than 15 half-lives). Simultaneously during the wait, since the half-life of 218Po is 3.071 minutes, 218Po comes to secular equilibrium with 222Rn during five half-lives. Then, 222Rn is measured for 30 minutes. If the anticipated 222Rn concentration was low (as with reference measurements), we used longer measurement times to attain a sufficient MDA.

Supplement 4 – soil gas measurements

A hardened Steel Soil Gas Probe from Durridge was pushed 25-35 cm into the soil and connected to the instrument via a vacuum gauge with water shutoff valve and 3-meter tubing. Soil gas was extracted from the ground using the instruments pump. The combined volume of the steel probe, water shutoff valve and tube volume combined are 0,115 litres. For corrections due to decay of 220Rn, please see Supplement 1 (Equation S.1 for RTM1688 since it is calibrated at the same flowrate and Equation S.1 and S.5 for EQF due to the flow rate of 0,75 l/min). On one day the EQF was the only instrument and on three days the RTM was the only applied instrument for 220Rn soil gas. On five field days, both EQF and RTM were used for soil gas measurements. This revealed a bias in 220Rn between the instruments where EQF254 had a mean of 6,300,000 Bq m-3 (SD: 1,400,000) while the RTM119 had mean of 4,400,000 Bq m-3 (SD: 1,000,000). The EQF was always used first and the RTM afterwards, which may have affected the bias. To account for the bias, and since EQF was used for all exhalation measurements, we used the mean of the bias to estimate what would have been EQF values on the three days when only RTM was used.

Supplement 5 – residual and QQ plots of statistical models

Mod 1 AIC=382 R2adj=0.64 F-stat: 24.39 on 3 and 37 DF,  p-value: 6.999e-09

Mod2 AIC=397  R2adj=0.70 F-stat: 24.56 on 4 and 36 DF,  p-value: 7.289e-10

Mod3 AIC=398  R2adj=0.64 F-stat: 36.57 on 2 and 38 DF,  p-value: 1.395e-09

Mod 4 AIC=355  R2adj=0.72 F-stat: 17.82 on 5 and 28 DF,  p-value: 6.142e-08

Mod5 AIC=331  R2adj=0.76 F-stat: 23.01 on 4 and 24 DF,  p-value: 6.457e-08

Mod6 AIC=436  R2adj=0.74 F-stat: 23.81 on 5 and 35 DF,  p-value: 2.315e-10

References:

Durridge Company Inc (2023) RAD7 Electronic Radon Detector User Manual, revision 2023-03-27 https://durridge.com/documentation/RAD7%20Manual.pdf.

Laraweb (2024) Laboratoire National Henri Becquerel" sitt "Library for gamma and alpha emissions.

Accessed 29 May 2024